Critical Hardware and Firmware Migration Challenges

Migrating from the controlled, validated environment of the EVM430-FR6043 to a miniaturized custom PCB introduces complex physical and electrical variables. The transition revealed four critical failure points across the acoustic, electrical, and firmware domains that had to be systematically isolated and resolved.

1. Acoustic Impedance & Signal Loss (The Transducer Interface)

- The Initial Behaviour: The 300kHz ultrasonic transducers functioned perfectly during open-air bench testing. However, once plastic spirometer flow pipe was mounted, the system suffered massive signal degradation.

- The Physics of the Failure: This failure was caused by a severe acoustic impedance mismatch. High-frequency acoustic waves (like 300kHz) have shorter wavelengths, making them highly susceptible to attenuation and scattering when encountering dense boundaries like plastic pipe walls. Instead of propagating cleanly through the air column, the 300kHz ultrasonic energy was being absorbed by the plastic or reflecting erratically, causing unpredictable resonance shifts and a distorted waveform envelope that the ADC could not accurately cross-correlate.

- The Resolution: Dropping the frequency to the Jiakang 200kHz transducers increased the acoustic wavelength, significantly improving the signal's ability to penetrate the plastic boundaries with minimal attenuation, thereby restoring a clean, readable Time-of-Flight envelope.

2. Clock Source Transition (EVM Ceramic Resonator vs. Custom Board Crystal Oscillator)

- The Hardware Baseline: The EVM430-FR6043 evaluation module utilizes a ceramic resonator for its clock source, which starts up very quickly but has a looser frequency tolerance. To improve upon this design, the miniaturized custom PCB was engineered from the start with a high-precision crystal oscillator to drive the main microcontroller clock.

- The Physics of the Challenge: The Ultrasonic Sensing Solution (USS) module utilizes a high-speed sigma-delta ADC to digitize microvolt-level acoustic waves at 1000 kHz, requiring an ultra-stable, low-jitter time base. Crystal oscillators provide pristine timing precision (in the tens of ppm) compared to ceramic resonators (around 0.5% or 5,000 ppm), which makes the crystal ideal for the USS cross-correlation algorithms. However, physical crystal oscillators rely on mechanical vibrations that take significantly longer to build up to a steady-state amplitude than ceramic resonators.

- The Firmware Mismatch: When the standard firmware—which was originally tuned for the EVM's fast-starting ceramic resonator—was flashed to the custom PCB, it expected the clock to be ready almost instantly. The MCU was commanding the USS module to fire the transducers and sample the ADC while the custom board's superior crystal oscillator was still physically ramping up, causing the ADC capture windows to drift and corrupting the telemetry.

- The Resolution: Recognizing the difference in startup characteristics between the two clock types led directly to the subsequent fixes: tuning the oscillator tank circuit with the correct 33 pF load capacitors and drastically increasing the firmware's crystal settling time to 5000 μs to accommodate the crystal's slower mechanical ramp-up.

3. Capacitance Mismatch (The Oscillator Tank Circuit)

- The Initial Behaviour: Even after upgrading to the crystal oscillator, the system clock struggled to start up reliably or maintain a stable oscillation.

- The Physics of the Failure: A crystal oscillator requires a very specific load capacitance (CL)to resonate at its target frequency. The custom board was initially populated with 12 pF load capacitors. This value was electrically incompatible with the requirements of the chosen crystal. Insufficient load capacitance prevents the internal oscillator amplifier from achieving the necessary loop gain to start the electromechanical vibrations of the crystal, or it pulls the frequency out of specification.

- The Resolution: The 12 pF capacitors were swapped out for 33 pF capacitors, matching the crystal's datasheet requirements and creating a tuned tank circuit that guaranteed robust, consistent oscillator startup upon power delivery.

4. Firmware Timing Prematurity (Crystal Settling Time)

- The Initial Behaviour: With the hardware fixed, the MCU would occasionally boot and attempt to capture ultrasonic data, but the initial telemetry batches were entirely corrupted before eventually stabilizing.

- The Physics of the Failure: Unlike internal RC oscillators that start almost instantly, physical crystal oscillators rely on mechanical vibrations that take time to build up to a steady-state amplitude. The firmware’s default HFXT (High-Frequency Crystal) settling time was set to a mere 120 μs. This meant the firmware was commanding the USS module to fire the transducers and sample the ADC while the crystal clock was still physically ramping up and highly unstable.

- The Resolution: The firmware source code was modified to increase the crystal settling time delay from 120 μs to 5000 μs (5 milliseconds). This massive increase acts as a necessary buffer, ensuring the microcontroller completely halts execution until the crystal's clock signal is perfectly clean, stable, and locked before initiating any acoustic captures.