## Slide 1
Standardization of Spirometry
ATS/ERS 2019 Update — Key Technical Standards & Clinical Implications
Sasanka Barman
Spirometer / Ultrasonic Flow Measurement System — MSP430FR6043
ATS/ERS Technical Statement • Am J Respir Crit Care Med 200:e70–e88 (2019)

## Slide 2
Agenda
01 Introduction & Scope
02 Equipment & QA
03 FVC/FEV₁ Maneuver
04 Acceptability Criteria
05 Key Parameters Reported
06 Grading System
07 Bronchodilator Testing
08 Device Design Implications
Why This Matters
Spirometry is the most common pulmonary function test. Accurate, reproducible measurements are critical for diagnosis, monitoring, and device development. The 2019 ATS/ERS update revised every major criterion — from maneuver acceptability to equipment accuracy to grading systems — impacting both clinical labs and embedded firmware developers.
2019
ATS/ERS Update
190+
Directly Relevant Studies
1,760
Patients Surveyed

## Slide 3
Standardization of Spirometry — Overview Mind Map

## Slide 4
Introduction & Scope
Spirometry — a physiological test measuring the maximal volume of air an individual can inspire and expire with maximal effort. The primary signal is either volume or flow as a function of time.
Primary Variables
FVC
Forced Vital Capacity — volume expelled from full inspiration with maximal effort
FEV₁
Forced Expiratory Volume in 1 second
FEV₁/FVC
Ratio indicating airflow obstruction
PEF
Peak Expiratory Flow
FEF₂₅₋₇₅
Mean flow in mid-expiratory phase
Document Scope
Definitions & terminology
Equipment specifications (ISO 26782)
Patient preparation & contraindications
Maneuver acceptability & repeatability
Quality grading system
Data reporting standards

## Slide 5
Equipment & Quality Assurance
ISO 26782 Compliance
▸ Max permissible error: ±2.5%
▸ Flow range: 0–14 L/s
▸ Volume range: ≥8 L (BTPS)
▸ Sampling rate: >100 Hz
▸ Min resolution: 12 bits
▸ Resistance: <1.5 cmH₂O·L⁻¹·s⁻¹
Daily QA Checklist
▸ 3-L syringe calibration verification
▸ Flow range: 0.5 – 12 L/s
▸ Accuracy tolerance: ±3%
▸ Zero-flow baseline set
▸ Filter used in test = filter in cal
▸ Log all cal results & failures
BTPS Correction
▸ All results at BTPS conditions
▸ Temperature accuracy: ±1°C
▸ For volume spirometers: measure internal temp per breath
▸ Error up to 6% if ambient temp assumed
▸ Alert if rapid changes >3°C in 30 min
3-L syringe accuracy: ±0.015 L or ±0.5% of full scale | Monthly leak test required | Dropped/damaged syringe = out of calibration until verified

## Slide 6
FVC / FEV₁ Maneuver — Four Phases
01
Maximal Inspiration
Rapid full inspiration to TLC. Pause at TLC <2 s. Slow or prolonged pause reduces PEF and FEV₁.
FIVC recorded here
02
"Blast" of Expiration
Immediate maximal forced expiratory effort. BEV must be <5% of FVC or 0.100 L. Rise time 10→90% PEF: <150 ms.
Back-extrapolation → Time 0
03
Continued Full Expiration
Sustain effort up to 15 s. EOFE plateau: <0.025 L change in 1 s. System beeps: 1 beep = plateau, 2 beeps = 15 s.
FVC, FEV₁, FEF₂₅₋₇₅ measured
04
Maximal Inspiration (Post)
Rapid return to TLC after EOFE. Provides FIVC — used to confirm starting volume. FIVC − FVC > 0.100 L → invalid.
Flow-volume loop completed

## Slide 7
Within-Maneuver Acceptability Criteria
Criterion
FEV₁ Accept.
FVC Accept.
FEV₁ Usable
FVC Usable
BEV < 5% of FVC or 0.100 L
✓
✓
✓
✓
No faulty zero-flow setting
✓
✓
✓
✓
No cough in first 1 s
✓
—
✓
—
No glottic closure in first 1 s
✓
✓
✓
✓
No glottic closure after 1 s
—
✓
—
—
EOFE plateau <0.025 L in 1 s
—
✓
—
—
No obstructed mouthpiece
✓
✓
—
—
No leak
✓
✓
—
—
FIVC − FVC < 0.100 L or 5% FVC
✓
✓
—
—
BEV = Back-Extrapolated Volume | EOFE = End of Forced Expiration | FIVC = Forced Inspiratory Vital Capacity

## Slide 8
EOFE Criteria & Between-Maneuver Repeatability
EOFE Acceptability (any one of 3):
①
Plateau
<0.025 L change for ≥1 s
→ Single beep
②
Time limit
FET ≥ 15 seconds
→ Double beep
③
Restrictive pattern
FVC ≥ prior FVC or within repeatability tolerance
(e.g. children, restrictive disease)
Between-Maneuver Repeatability
Metric
Age > 6 yr
Age ≤ 6 yr
FVC diff (2 largest)
< 0.150 L
< 0.100 L or 10%
FEV₁ diff (2 largest)
< 0.150 L
< 0.100 L or 10%
Min acceptable maneuvers
3
3
Max maneuvers (adults)
8
> 8 possible
Safety threshold: If FEV₁ from acceptable maneuver drops below 80% of start value, terminate testing.
Goal: ≥3 acceptable FEV₁ AND ≥3 acceptable FVC (may come from different maneuvers)

## Slide 9
Quality Grading System (ATS 2019)
* Or 10% of highest value, whichever is greater — applies for age ≤ 6 years only. Graded separately for FEV₁ and FVC, and separately for pre- and post-bronchodilator sets.
Grade
# Maneuvers
Repeatability (>6 yr)
Repeatability (≤6 yr)
A
≥3 acceptable
< 0.150 L
< 0.100 L*
B
2 acceptable
< 0.150 L
< 0.100 L*
C
≥2 acceptable
< 0.200 L
< 0.150 L*
D
≥2 acceptable
< 0.250 L
< 0.200 L*
E
≥2 acceptable
> 0.250 L or 1 accept.
> 0.200 L*
U
0 acceptable, ≥1 usable
N/A
N/A
F
0 acceptable, 0 usable
N/A
N/A
Key Points
• Grade A = highest confidence result is patient's true best
• Grade U = usable but not acceptable
• Grades < A are still clinically informative
• Grading informs interpreter confidence, not lung health
• Goal is always best quality per patient session

## Slide 10
Bronchodilator Responsiveness Testing
Protocol Flow
01
Baseline
Spirometry
3 acceptable
FEV₁ & FVC pre-BD
02
BD
Admin.
Dose/route per
facility protocol
03
Wait
Period
Per agent
withholding time
04
Post-BD
Spirometry
3+ acceptable
FEV₁ & FVC
Bronchodilator Withholding Times
Agent Class
Example
Withhold
SABA
Albuterol / Salbutamol
4–6 h
SAMA
Ipratropium bromide
12 h
LABA
Formoterol / Salmeterol
24 h
Ultra-LABA
Indacaterol, Vilanterol
36 h
LAMA
Tiotropium, Umeclidinium
36–48 h
Reported Values
FVC
L (2 decimal places, BTPS)
FEV₁
L (2 decimal places, BTPS)
FEV₁/FVC
Decimal fraction to 0.01
PEF
L/s
FET
Seconds
FIVC
Largest inspiratory VC post-EOFE
Δ FEV₁ / FVC
Absolute + % change vs. pre-BD
FEF₂₅₋₇₅
From maneuver: largest FEV₁+FVC sum

## Slide 11
Device Design — ATS/ERS Requirements for Developers
Signal Acquisition
▸ Sampling rate > 100 Hz
▸ Min ADC resolution: 12 bits
▸ Flow range: 0 – 14 L/s
▸ Volume: 0.5 – 8 L (BTPS)
▸ Back-extrapolation for Time 0
▸ BEV detection & flagging
Real-Time Processing
▸ Simultaneous flow-volume & volume-time display
▸ PEF rise time detection (<150 ms threshold)
▸ Plateau detection: <0.025 L in 1 s
▸ EOFE audio alerts (1 beep/2 beep)
▸ FET timer with 15 s hard stop
▸ BTPS compensation per breath
QA & Reporting
▸ Automated acceptability feedback
▸ Grade A–F + U flagging
▸ Calibration log generation
▸ ±2 SD alert on cal factor drift
▸ Zero-flow verification before each test
▸ Operator override capability
MSP430FR6043 context: USS SW Library computes AbsTOF_UPS/DNS → VFR at 400 Hz → all 6 requirements above directly map to firmware pipeline design

## Slide 12
Key Takeaways
Equipment
ISO 26782 compliance; ±2.5% max error; >100 Hz sampling; daily 3-L syringe cal verification
Maneuver
4-phase FVC; BEV <0.100 L; EOFE plateau <0.025 L in 1 s; back-extrapolation for Time 0
Acceptability
9 within-maneuver criteria; min 3 acceptable for each of FEV₁ & FVC; ≤8 adult attempts
Grading A–F, U
A = highest confidence; repeatability tolerances tighten for younger patients; graded per set
BD Testing
Pre & post BD spirometry; agent-specific withholding (4 h – 48 h); report absolute & % Δ
Device Design
>100 Hz ADC, real-time BTPS, plateau detection, dual-display, automated operator feedback required
ATS/ERS 2019 — Miller et al., Eur Respir J 2005; Graham et al., Am J Respir Crit Care Med 2019